## Slide 1: K factor

## Slide 2: My Initial Set-Up (TG-51)
80 cm
10 cm
10 cm
Replicated the TG-51 geometry for Co-60 in MC, took voxels approx. equal to cavity radius (0.305)
Ran sim x 10 -> 2.0892E-15 Gy/Particle
Took Aug 2020 source abs dosimetry
k = 9.76 E 14
Dose Rate (SSD =80, 10x10, d=10) = 203.61 cGy/min

## Slide 3: Levi’s TGI-51 Measurement
Single run, single voxel
K = 8.78 E14 (smaller than mine)

## Slide 4: Levi’s Extended Sweeping w/ Flattening Filter
Using TGI-51 vs “Treatment” rdg ratio to convert dose rate (2020) from TG51 geometery to Treatment
Dose rate = 11.825 cGy/min
Levi modelled sweeping (10 sweeps)
k = 9.94 E 14 (slightly larger than my value)

## Slide 5: Supporting evidence to increased particles with treatment geometry

## Slide 6: Other numbers on the wind
K = 1.17 E 15 (thesis) “Treatment Setup” but 5 sweeps?
K = 1.1531 E 15 (instructions)

## Slide 7: Options
Use my existing value (9.76E14)
Use the measured dose rates from LB and run my own sweeping sim to recompute (?)
Use Levi’s sweeping value (9.94E14)
Tune the value to achieve Rx at Rx point (?)