TBI Monte Carlo: procedure

Refer to “detailed procedure” for screenshots and some debugging discussions

For reference, the student account to the cluster has password: stu!_D@2

The codes are now all on the cluster.

Start by copying all of the codes into a new folder for the patient (in the Unix shell, not on Windows).

E.g. folder should be called “12345678_AP”

The codes are here:

/home/tbi_clin/MC_codes/codes

(use command cp * /path/to/destination/, the * selects all files in folder)

Just in case you lose them, the codes are also located here:

/home/tbi_clin/MC_codes/source_codes (here they are with the .cc codes that create the executables above)

And also:

H:\Physics\VCC\Treatment Techniques\TBI\Levi Monte Carlo\source_codes

The files that we need:

- The patient egsphant from the tbi_clin account (e.g. 1206501_AP.egsphant)

  - Make a new folder for each simulation (i.e. one for the patient prone, one for supine)

- The contourPoints file from the same folder (e.g. LOG.contourPoints or 1111.contourPoints)

  - Put this in the same folder as the Python code

- The Excel spreadsheet for the patient treatment

- The package of C++ scripts; only need to be recompiled if you changed the .cc file for any reason

We need to get some numbers from the egsphant – t0 coordinates, bed_y, and the grid dimensions

- Use dosxyz_show to see the x, y, z coordinates of t0 and bed_y

  - t0 is the treatment marker at the umbilicus (xz plane)

  - bed_y is the y-coordinate of the lowest part of the patient (look at xy plane)

  - Command is: dosxyz_show 1206501_AP.egsphant

  - Note: the display always says “x = 0.000, y = 0.000”, but it’s only actually x and y if you are on the “xy” grid display. If you are clicking yz, xz, then it actually means these variables.

- Get grid dimensions: Use the code “read_out_vals.cc”, compiled to “read”, with this syntax:

  - ./read 1206501_AP     (do not include the extension “.egsphant” here)

- From the Excel patient sheet, find the following items

  - MC_scaled z-values for t0, t1, t2.

  - Treatment time

  - Lung compensator thickness

  - Date of first treatment

  - Number of styros

First: type ./shift_calcs and press enter. Then follow prompts for entering bed_y, t0_y (as a negative number), number of styros, and plastic thickness (normally 0.6) as prompted. It will output shift_y_plastic, shift_y_lead, and SSD for the following calculations.

We have to prepare .egsphant files for the lung compensators and the plastic tray. These need to have the right dimensions and the right geometry. We use quick scripts in python and Matlab for this:

Python

Make sure that the python code, read_and_scale_CPs_args.py, is in the same folder as the LOG.contourPoints file.

If the contourPoints file is named anything else, change the name of the file by using:

mv oldfilename.contourPoints LOG.contourPoints

Type this into the command line on the cluster:

python read_and_scale_CPs_args.py t0z t1z t2z (Use at least one decimal, including 0.00 for t0z)

These three values are the MC_Scaled values from the Excel sheet.

After running this Python script – you now have a .txt file, called contourstomask_res.txt

It will give you some warnings about some functions being deprecated. Don’t worry about them.

Matlab: (to open, type “matlab” in the cluster)

NOTE: This must be done on the student account on the cluster (due to read/write privileges).

Username: student	 password: stu!_D@2

Use the code corresponding to the phantom that you’re making (AP, Supine or PA, Prone):

mask_from_contours_2_AP.m

mask_from_contours_2_PA.m

Lines to change: from line 35 to line 45.

- xmin, xmax, zmin, zmax

  - Get these values by opening the patient .egsphant and looking at the coordinate grid at the start – first and last values of the x-row and z-row

  - Enter them all as POSITIVE values!

  - From the “read” code

- xdim, zdim

  - The dimensions of the phantom (from the very first line of the egsphant; the first and third values)

  - Also from the “read” code

- t2z_scaled: same as the value entered in Python

- compthick: the thickness of the compensator

- tray thickness (e.g. 0.6)

After you run these scripts, Matlab will open a window with a picture of the compensators (they will be upside-down). If they are blank, it’s because the L and R lungs were contoured in a different order by the rad onc (my code follows what they normally do, which is L then R)

If this happens, all you need to do is open the .txt file that Python created (contourstomask_res.txt), and cut and paste the first two lines to the end of the file. This tells Matlab that the lungs are in the other order. Save it and run the Matlab script again.

- Move contourstomask_res.txt to the original folder

  - Needs to be copied/moved using the tbi_clin account, otherwise it says “Permission denied”

By now, confirm that you have three files in your working directory: the patient egsphant (1206501_AP.egsphant), the plastic tray (plastic3.egsphant), lungcomps.egsphant.

C++:

Creating the complete egsphant file (patient + tray + comps):

Each of the codes produces a new .egsphant. You’ll want to delete the intermediate .egsphants when you finish with a patient in order to clear disk space.

If you change the codes, you need to recompile them and make new executables. I’ve explained how to do this in the longer document. The source codes are all in this directory:

/home/tbi_clin/MC_codes/source_codes

Put the plastic tray, lungs on top of the patient, and then extend to make room for flattening filter:

./puts_extend 1206501 AP shift_y_plastic shift_y_lead extend_pixel

(extend_pixel is the number of pixels we want to extend in the inferior direction 1 cm = 4 [for a resolution of 0.25 cm])

Putting the plastic tray on top of the patient:

./plastic 1224849 AP shift_y_plastic

Putting the lead compensators on top of the plastic tray:

./lead 1206501 AP shift_y_lead

Extend the patient phantom, to make room for the entire flattening filter:

./extend 1206501 AP

Shift the filter egsphant over t0 and then stack on top of the patient phantom:

./shift_stack 1206501 AP t0_x t0_y t0_z SSD (enter at least one decimal for each)

./shift_stack 1206501 AP t0_x t0_z SSD t0_y (enter at least one decimal for each)

./filtershift 1206501 AP Tx_x Tx_z SSD Tx_y

./stacker 1206501 AP Supine

The output is called 1205601_AP_complete.egsphant

- These steps need to be iterated a few times (this code was originally designed for phantoms of 0.5 cm resolution).

- Use dosxyz_show to look at the “*_complete” egsphant. Go to the bottom slice of the filter (around slice 21) and look at the x, z boundaries of the filter. Take the difference between the boundary values and add these values to the t0_x, t0_z values.

  - E.g. If you originally used t0_x = 0.1 cm, t0_z = -0.2 cm, and then you find the boundaries to be (-10cm, 20cm) for x and (-15cm, 45cm) for z, then your new t0_x value is 0.1 + (20 – 10) = 10.1 and your new t0_z value is -0.2 + (45 – 15) = 29.8.

  - After this, if you’re still off by more than 0.5 cm in either direction, manually repeat filtershift/stack with small adjustments.

  - If you have to iterate multiple times, (e.g. a 0.5 shift in x-direction only), apply this shift to the previous values, don’t restart from zero (e.g. if you have a 13.0 x-shift and you need to shift 0.5 more, enter 13.5 (the total), not just 0.5).

**Important: In order for the MC simulation to work, you need to open the egsphant and delete the duplicate material listed.

- emacs *_complete.egsphant is how to open the editor

- Reduce the number of materials (first number) from 10 to 9

- Delete the second PMMA material entry

- Delete one of the zeroes in the following line (there should be one zero for each material)

The final egsphant will be called 1206501_AP_complete.egsphant

Creating the egsinp file:

Edit Create_egsinp.cc to enter the SSD, coordinates of t0 (in here the variables are called  Cx/Cy/Cz) and the output file name. Everything else, you can edit after the .egsinp file is created.

The egsphant can stay in the folder you have been working in. The egsinp file must go to dosxyznrc. Make sure the egsinp folder in dosxyznrc points to the right directory before you start the simulation!

Use emacs Create_egsinp.cc to edit

g++ -Wall Create_egsinp.cc –o create

./create

Now, put the .egsinp and .egsphant file in the MC cluster.

The .egsinp should have the same name as the .egsphant (1206501_AP_complete)

Simulation:

- Move the .egsinp and .egsphant file to the EGS_HOME/dosxyznrc directory in tbi_clin.

- Open the .egsinp file and replace the file destination for the .egsphant with the right file location Run the simulation:

exb dosxyznrc 1206501_AP_complete 700icru_TT long batch=at p=256

- I normally check (condor_q) to make sure there is a _w1 file after the simulation starts running, or else you know already that the simulation is not going to complete properly (if you forget to decrement the medium number then this will happen)

- After the simulation

  - Retrieve the .3ddose file, put it in the folder you have been working on so far.

  - delete all of the *_w* files to clean up dosxyznrc

  - The .3ddose file will have the same name as the egsphant and egsinp (e.g. 1206501_AP_complete.3ddose)

Processing the dose files (C++ again):

./strip 1206501_AP_complete

./sgolay 1206501_AP_complete_trimmed.3ddose 3

- Need to calculate the conversation factor (I have an Excel sheet)

Formula: exp⁡(-ln2days from Aug 13 2015 to Tx start1921.5 days)*(treatment time per field)*6*1.1849*1015

(Assuming it’s a 12 Gy/6 fraction treatment, that is, 6 Gy for each Supine and Prone files)

Then put this in line ~295, needs to be hard-coded right now

./conv sgolay_1206501_AP_complete_trimmed conv_factor

(The conv factor is hard-coded to be some value times 10^16, you don’t need to put the 10^16 in.)

NB. If you are revisiting this sometime later than late 2020, there will be a new Cobalt source. Also applies if you ever want to simulate patients from before 2015

I made this measurement based on the activity of the Aug 13 2015 Cobalt source on the day it was calibrated.

It’s not hard to correct for a new source. In an Excel sheet, go to the third page “Internal info”, cell B12 is the “Cal. Output: 10x10 @ 80 cm”. (373.74 for patients from late 2010 to mid-2015). For any MC simulations with a new Cobalt source, just multiply the calibration factor that I have by a ratio, (NEW CAL OUTPUT)/(373.74)

Eclipse/MIM:

We need to put the patient from Eclipse into MIM (twice; once as Supine, once as Prone), anonymize, and bring them back into Eclipse. Will be treated as two separate patients.

- For each treatment position: import the CT image volume into MIM from Eclipse

- The CT image volume will appear in the “INCOMING” folder on MIM, with the patient’s ID

- Move this CT image to the “TBI_PROJECT” folder (use the “Send To” button)

  - Right-click on the image to find the Anonymize option.

  - I used the name convention “VC_TBI01S” (Vancouver Centre, two-digit patient ID – I kept a master list of mapping my IDs to the real IDs as per ethics rules. S for Supine, P for Prone)

- Send VC_TBI01S/P back to Eclipse (Export and re-import)

  - After they’re anonymized you should delete the original patient files in INCOMING to clear up space

- In Eclipse:

  - After the patient is brought back, make some field and calculate the doses

    - Make sure you are getting ABSOLUTE dose, not relative. This option is only available if you choose some reference point for the dose calculation (it doesn’t matter which you pick; the dose file here is only being provided so that we can overwrite it with the DICOM tags)

    - I normally made some 35x35 cm^2 field, this was just a habit, doesn’t matter.

  - Export the dose (RD file)

- In MIM: Contouring

  - Lungs: do L and R separately, I used the seed fill feature

  - Kidneys (do L and R separately), liver: do every 2-3 slices and then auto-interpolate between the slices, then go through with the brush to polish it off a bit

  - Thyroid: can be harder, do every slice on its own (sometimes it only spans over 7-8 slices)

  - Body: use the Whole Body feature (need to use Sagittal or Coronal views for this tool) and then remove top 4 and bottom 8 slices, then use brush (3D brush) to remove the arms.

DICOM:

Now we make a DICOM from this dose file. This is where we overwrite that RD file we pulled from Eclipse before.

python Create_RD_DCM_debug_TT_UCSD.py RD_from_Eclipse_name.dcm 3ddose_file_name.3ddose /path/to/this/folder/ New_RD_name.dcm

You can type “pwd” in the command window to get the path.

Put this new .dcm file somewhere in the H drive where you can then upload it to MIM (put it in a new folder on its own; in MIM you can only select folders, and then it picks the dose file from the folder).

Contour from there, I normally did this on the Supine image.

*All of the above is repeated for Supine and Prone.*

MIM

Importing the dose file into MIM is easy, but make sure that you make a new folder just for the one dose file you are interested in. In the MIM Import setting, you can select folders, and then inside that folder, it will pick out the dose file and import it.

Be careful to import to “TBI_PROJECT”. If you import it somewhere else, either re-import it to the right folder, or go find it in its destination and click “Send To…” to send it to TBI_PROJECT.

Prone dose issue in MIM

At this time, there is some sort of disagreement between the above DICOM file and the way that it is imported into MIM, only for Prone images. If some people who know a lot about DICOM files sit down and figure this out, you may be able to get rid of this step…

When you bring the prone dose file into MIM, you’ll find that it is translated in the x (left-right) and z (up-down) directions away from the actual CT image. You’ll need to measure this offset and create a new 3ddose file to manually bring the dose back to where it needs to be.

- Use the measure tool in MIM to measure the offset in x and z (doesn’t need to be perfect for the first iteration)

- Using the x-bounds and z-bounds from your data sheet, add the value of the offset to x and z to determine what your new grid dimensions have to be.

  - E.g. if your x-offset was 10.0 cm and your current x-grid spans (-20.0, 20.0), your new x-grid should span (-10.0, 30.0).

- When you decide the new values, use ./arrayvals to produce a .txt file with the new grid in it. (Do this once for x and once for z.)

- Open these new .txt files, as well as the final .3ddose file that you previously converted into a DICOM file. Replace the x and z grids from the .3ddose file with the ones from the .txt files

  - Save this .3ddose file as a new file (e.g. just add “_shift1” to the end of the filename)

  - Placing an ampersand & after the emacs command lets you have several text windows open at a time (e.g. emacs filehere.txt &)

- Repeat the above DICOM step with the new dose file (same initial RD file)

  - Make sure to enter the text of the new .3ddose file

- Bring this new RD file back to MIM. Hopefully the offset is a lot better. Use the measure tool again and iterate again until the prone dose is overlaid on the prone CT.

  - Note: Occasionally, when you are almost at the correct position and increment one more centimeter, it will all of a sudden jump 30-40 cm away… if this happens, just keep iterating and add 30-40 cm to that grid again. Not sure why this happens.

A note about file transfer on the upstairs computers: I normally use both the MIM computer and the computer two to the left of it (beside the door) for the above steps. To bring the .3ddose file to the H drive (or wherever else you are importing into MIM from), follow these steps:

- Open the Moba program on the non-MIM computer (if Moba can be installed on the MIM computer it will save a lot of energy)

- Open two windows. In one, sign into gundam. On the other, use “cd ..” to work your way to the “drives” folder, and go to the h drive, to wherever you are importing into MIM from.

- In the folder where you are in the h drive, type:

  - sftp tteke@205.233.175.86 (password tteke when prompted)

- Use “cd” commands to get to where, in gundam, your DICOM file is stored.

- Type “get filename.dcm”

  - (Can also do this in reverse order by sftping to the h drive from gundam, then use the “put” command)

- When you are done in “sftp” mode, type “bye” to return to the normal screen display

Deformable registration

Once you bring both the Supine and Prone dose distributions into MIM, control+click on the two CT scans and two corresponding dose files (a total of four files), and then double-click the “TBI Accumulation” button on the right-hand side of the screen. (You’ll have to select the “Workflows” tab on the right-hand side of the screen to get this option available.) Nevin made this button.

Every time it pauses or prompts you to do something, just tell it to keep going (click Continue). There are opportunities to use RegReveal and RegRefine to make the registration better but I have not used this myself.